## Slide 1
LED calibration pulse measurement
Calibration pulse generation circuit
FPGA_DAC
6-bit
500 MSPS
DAC_D(5)
DAC_D(4)
DAC_D(3)
DAC_D(2)
DAC_D(0)
DAC_D(1)
DAC_OUT
FPGA
1, 0, Z
DAC_Zero
LED Cal pulse
2 op-amp in parallel drive 36-channels
To another four sets
FPGA_enable(1)
FPGA_enable(36)
FPGA_enable(…)
FPGA_enable(2)
LED Cal pulse
LED(1)
LED(36)
LED(…)
LED(2)
Ground

## Slide 2
Calibration pulse vs DAC setting (oscilloscope picture)
DAC setting,
0x7f
0
Width
Width
4ns
12ns
16ns
20ns
24ns
28ns
32ns
8ns
Trigger output
Calibration
pulse

## Slide 3
Calibration pulse vs DAC setting (oscilloscope picture)
Trigger output
Calibration
pulse
Zoom in on the smaller pulse
Edge slop:
~1V/6ns

## Slide 4
The measurement at the BLUE LED is unreliable, but here is an example with floating probe ground
Trigger output
Probe at the VLD board
Probe at the blue LED anode (not the same LED as NPS), The LED cathode back to VLD (another ~150 ft wire/cable)

## Slide 5
Set the calibration pulse baseline to about 2V to compensate for the BLUE LED light emission threshold
Trigger output
Probe at the VLD board
Probe at the blue LED anode (carhode to ground

## Slide 6
LED calibration pulse measurement
Calibration pulse generation circuit
FPGA_DAC
6-bit
500 MSPS
DAC_D(5)
DAC_D(4)
DAC_D(3)
DAC_D(2)
DAC_D(0)
DAC_D(1)
DAC_OUT
FPGA
1, 0, Z
DAC_Zero
LED Cal pulse
2 op-amp in parallel drive 36-channels
To another four sets
FPGA_enable(1)
FPGA_enable(36)
FPGA_enable(…)
FPGA_enable(2)
LED Cal pulse
LED(1)
LED(36)
LED(…)
LED(2)
Ground
A
E
B’
B
C
F
D
The following (until next date mark, on page 21) are measured on Feb. 28th, 2022

## Slide 7
Calib at the cable (30 cm from PCB)
The ‘DAC’ + input
A
B
The ‘op-amp’ end of DAC output

## Slide 8
The ‘op-amp’ + input
C
D
The ‘op-amp’ output

## Slide 9
The same, but 2V vertical scale, to compare with Calib on the cable
E
E
The ‘op-amp’ after the 0.5Ohm

## Slide 10
Disable the other op-amp
E
B’
Probe at the DAC- input

## Slide 11
Probe at the DAC output
DAC-
DAC_o
Probe at the DAC output with one op-amp

## Slide 12
DAC output with one op-amp active
B’
Only one CH
Op-amp output, with all the other op-amp disabled

## Slide 13
Same as before, after 0.5 Ohm
No_amp
36-ch
With all the 36 switches ON. Lower amplitude and slower calib pulse

## Slide 14
DAC output with one Ch (one switch on)
One opamp
No opamp
DAC output with zero Ch (all the switches are OFF)

## Slide 15
LEMO output, 4ns, (100mV/grid)
0x1
0x2
LEMO output, 8ns

## Slide 16
LEMO output, pulse width 32ns
0x8
0x8
Accumolated LEMO, pulse width 32ns

## Slide 17
Reproduce the pulse shape
Remove the DAC compensation cap (5pF), the DAC output is oscillating at ~50MHz. This plot is the average (to average out the osc.), the actual DAC output is faster

## Slide 18
Add a 5pF cap in parallel with RFP13/18, the NIM output is better
Add the Cc back to the DAC, Remove one Cc of the two Calib opamp. The average pulse
Same as left, except it is a single trigger, oscillation, and faster calib signal

## Slide 19
Remove both Cc, the oscillation is ~35 MHz, though the Calib pulse is faster
Add a 5pF in parallel with RFP16, no visible change
Replace 3.3pF compensation capacitors. It is stable, and the pulse is bigger. (slop is faster)

## Slide 20
Larger scale to show the falling edge

## Slide 21
When the probe touched pin#5 of MCH6541, the ECL output falling edge improves a lot. The green shows the pin#5, which should be a constant.
Add a 2.2pF cap from pin#5 to ground.
Also replace the 5pF Cc with 2.2pF for the DAC, and Top_onboard opamp.
The DAC output is oscillating (cold solder)
The following are measured on Mar. 2nd, 2022

## Slide 22
Reflow solder for Cc of the DAC, and also increase to 4.4pF (I think that it is still 2.2pF). The amp output is stable with 2.2pF compensation. This is the TOP_onboard connector (2.2pF)
This is the bottom_onboard connector. (the DAC Cc might be 4.4pF) (3.3pF Cc for the Calib Op-amp set on Feb. 28th) GREEN: DAC output, PURPLE: CALIB on the cable.

## Slide 23
Small pulse (4ns), the signal seen on the cable. It can be ~20ns, and up to 1.5V (above the LED threshold). This is the front panel bottom connector, 5pF compensation capacitor
The linear DAC setting should be:
0x00, 0x01, 0x02, …, 0x13, 0x14, 0x40, 0x41, 0x42, …, 0x7e, 0x7f (bit#6 is DAC_ZERO, the designed base line)
The DAC_ZERO is equivalent to 0x11 as seen on the setup,  1.64K.
The actual resistor is 1.33K on VLD.

## Slide 24
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 5pF, 4ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
1.3V, Base_Width: 8ns
1.3V, Base_Width: 12ns
1.0V, Base_Width: 13ns
1.4V, Base_Width: 28ns

## Slide 25
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 0pF (removed), 4ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
1.2 V, Base_Width: 8 ns
2.5 V osc, Base_Width: ? ns (50 MHz)
1.8Vosc + 1.0 V, Base_Width: 20 ns (50 MHz)
0.5Vosc + 1.8 V, Base_Width: 30 ns (50 MHz)

## Slide 26
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 2.2pF, 4ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
1.3 V, Base_Width: 8 ns
1.8 Vosc + 0.8V, Base_Width: 65 MHz
0.9 Vosc + 0.6V, Base_Width: 65 MHz
0.5 Vosc + 1.7V, Base_Width: 65 MHz

## Slide 27
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 4.4pF, 4ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
Vcal on OnBoard_top connector output:
1.2 V, Base_Width: 8 ns
1.3 V, Base_Width: 12 ns
1.0 V, Base_Width: 13 ns
1.3 V, Base_Width: 28 ns
1.4 V, Base_Width: 28 ns

## Slide 28
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 3.3pF, 4ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
Vcal on OnBoard_top connector output:
1.3 V, Base_Width: 8 ns
1.6 V, Base_Width: 12 ns
1.1 V, Base_Width: 14 ns
1.5 V, Base_Width: 28 ns
1.5 V, Base_Width: 26 ns

## Slide 29
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 3.3 pF, 8ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
Vcal on OnBoard_top connector output:
1.4 V, Base_Width: 8 ns + 4ns
2.3 V, Base_Width: 20 ns
1.7 V, Base_Width: 20 ns
2.5 V, Base_Width: 36 ns
2.6 V, Base_Width: 36 ns

## Slide 30
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 3.3 pF, 4ns pulse (0x11 as base, 0xFF square wave), added a 12pf capacitor over the 220 Ohm resistor of amplifier input (between DAC and Opamp)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
Vcal on OnBoard_top connector output:
1.2 V, Base_Width: 8 ns
1.5 V, Base_Width: 12 ns
1.1 V, Base_Width: 14 ns
1.5 V, Base_Width: 28ns
2.5 V, 36ns (input: 1.6V, 20ns by 8ns pulse)

## Slide 31
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: 15pF, 4ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
1.2 V, Base_Width: 8 ns
0.6 V, Base_Width: 16 ns
0.5 V, Base_Width: 16 ns
1.1 V, Base_Width: 24 ns

## Slide 32
The following are measured on Mar. 21st, 2022, the DAC compensation capacitor study.
Compensation capacitor: pF, 8ns pulse (0x11 as base, 0xFF square wave)
Vin (near the DAC input pin):
Vout (near the ADCop-amp of the resistor):
Vop-amp (near the ADCop-amp input pin):
Vcal (on the long cable, near the VLD):
Vcal on OnBoard_top connector output:
1.2 V, Base_Width: 8 ns + 4ns
0.9 V, Base_Width: 20 ns
0.7 V, Base_Width: 22 ns
1.7 V, Base_Width: 30 ns
1.8 V, Base_Width: 30 ns

## Slide 33
The following are measured on Mar. 25th, 2022, Check the Calibration signal rising/Falling time with analog switch
Compensation capacitor: pF, 8ns pulse (0x11 as base, 0xFF square wave)
With ~5V pulse, the rising time is ~5ns,
The falling time is ~10ns,
The recovery time (after switched off) is ~us, which is caused by the large capacitance of the long cable
A faster calibration signal (smaller rising edge and falling edge) can be generated this way than the op-amp alone.

## Slide 34
The following are measured on Mar. 28th, 2022, Check the Calibration signal rising/Falling time with analog switch and calibration pulse setting coincidence (also with TRIGGER signal)
Set the Swith_Enable width to 0x77 (very large), and move the switch_enable toward the pulse. From left to right, the 0x74 register is set to: 0x7700, 0x7720, 0x7736, 0x7738. The pulse shape change might be caused by the cable (reflection).
The rising edge is ~5ns for a 8V pulse, much faster than the op-amp slop

## Slide 35
The following are measured on Mar. 28th, 2022, Check the Calibration signal rising/Falling time with analog switch and calibration pulse setting coincidence (also with TRIGGER signal)
For the same switch_Enable delay, change the swith_enable width. The register 0x74 is set (from left to right plots) to 0x0836, 0x0636, 0x0436, 0x0A36 (the time scale was changed from 20ns in plot#1 to 10ns in plot#2, 3, and 4)
The falling edge is much slower than the rising edge.

## Slide 36
The following are measured on Mar. 28th, 2022, Check the Calibration signal rising/Falling time with analog switch and calibration pulse setting coincidence (also with TRIGGER signal)
The following plots are some sample pulses (as seen on the cable) with various settings.
Initial pulse width at the DAC input: 44ns;
0x74 register setting: 0x0636
Initial pulse width at the DAC input: 52ns;
0x74 register setting: 0x0633
Initial pulse width at the DAC input: 80ns;
0x74 register setting: 0x1228