## Slide 1: vXS fPGA-based Time to Digital Converter (vfTDC)
Trigger Interface
Trg/Clk/Reset/Busy
4x fiber Tx/Rx
VXS P0:
Trg/Clk/Reset/Busy
32 LVTTL in
32 differential in
Generic
8 differential In
8 ECL out
32 LVTTL in (with mezz. for differential IN)
32 differential in
28 LVTTL on row A
4 + 4 LVTTL on row D
28 LVTTL on row C
VME64x:
Register,
Data Readout
FPGA, XC7A200T-2FF1156C
18Mbit RAM
preliminary

## Slide 2
FPGA-based TDC measurement:
It measures the edge relative to the clock directly,
  How many clock cycles passed (coarse measurement);
  Latency between the signal edge and the clock edge (fine measurement)
The common start (or common stop) can be used as a normal channel, (multiple measurement on the common signal can decrease the statistical measurement error)
If the clock is accelerator synchronized clock, there will be no need for the precise common start (or common stop). The readout trigger should be sufficient.
Easier PCB design. Actually, this is based on a PCB designed for HallA.

## Slide 3
Ch_In
Calib_In
Calib_sel
Carry Chain
32x4 units
Clock
500MHz
FD
Register
128
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
FD
Edge finder, rising/falling edge
Pipelined encoder
R_edge F_edge
Fine_Time(6:0)
18Kb ring buffer,
4 us deep
R_edge F_edge
Fine_Time(6:0)
Dly_R_edge Dly_F_edge
Dly_Fine_Time(6:0)
Clock
250 MHz
EdgeInfo(17:0)
Trigger
Readout Lookback
Readout window
Ring_Buf_read_ddr(9:0)
CoarseTime
FineTime
Event FIFO
36(wide) x 512(deep)
Event
marker
“ZERO”
suppressed
Single channel measurement:
(one edge per 2ns)

## Slide 4
Ch#1 data(26:0)
Ch#2 data(26:0)
Ch#8 data(26:0)
Merged data(26:0)
27(wide) x 2048(deep)
Merged data(26:0)
27(wide) x 2048(deep)
Merged data(26:0)
27(wide) x 2048(deep)
Merged data(26:0)
27(wide) x 2048(deep)
Further merged
data(26:0)
(connector#1)
27(wide) x 2048(deep)
Merged data(26:0)
27(wide) x 2048(deep)
Ch#32 data(26:0)
Ch#192 data(26:0)
Further merged
data(26:0)
(connector#6)
27(wide) x 2048(deep)
Further merged
data(26:0)
(All channels)
27(wide) x 4096(deep)
6  1 merger
Trigger/Clock/Reset
Event Headers
32(wide) x 1024(deep)
Full block
VME Readout
Other Readout
Block headers
event#1 headers
event#1 data
------
event#n data
Block Trailer
Data Format
Bits (31:27, 26-24, 23:19, 18, 17:8, 7, 6:0
Format Connector# Channel# EdgeType Coarse_time 2ns Fine_time
xxxxx 0 - 5 0 - 31 1: R 0: F 4*(0 – 1023) ns 0/1 0 - 127
8  1 merger
4  1 merger
TDC readout (data merging):

## Slide 5
Some edges will be measured twice. This gives a real-time carry chain delay calibration, esp. the fine delay step size (LSB);
The FPGA built in IODelay elements can be used to calibrate the delay chain linearity;
External calibration is needed only for the delay offset.
Programmable: by sacrificing the channel count (192), the vfTDC can expand its measurement ranges (4 us) (limited by the total available memory size); Choosing a higher performance FPGA, the measurement LSB can be improved.
Special features of the vfTDC:

## Slide 6: vfTDC in-FPGA calibration
Delay (~us)
Ch#1 TDC
VME_Calib
IODELAY
Initialize
Increment
Input
Output
31.25MHz
clock
Calib1
IN1
Calib32
Calib_Trg
IN32
TRG
Calib_Sel
Ch#32 TDC
Each group (32 channels) has one set of calibration logic, so the IODELAY is localized
One rising edge and one falling edge per 32ns.
IODELAY increases (2.5ns/32 ~78ps) every 32ns.

## Slide 7: A typical channel in-FPGA calibration
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 10 | 11 | 11 | 12 | 12 | 12 | 13 | 13 | 13 | 14 | 14 | 15 | 15 | 16 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
 | 189 | 197 | 199 | 201 | 207 | 212 | 215 | 221 | 225 | 231 | 233 | 238 | 258 | 243 | 261 | 247 | 265 | 251 | 269 | 253 | 273 | 279 | 283 | 287 | 293 | 295 | 299 | 299 | 301 | 305 | 309 | 314 | 319 | 325 | 327
 | 196 | 202 | 210 | 214 | 218 | 222 | 230 | 231 | 238 | 242 | 259 | 246 | 263 | 250 | 267 | 252 | 271 | 278 | 282 | 286 | 290 | 294 | 298 | 299 | 299 | 303 | 310 | 311 | 318 | 322 | 327 | 330 | 334 | 339
(In 78ps steps)
Rising edge
Falling edge
(TDC Count)
Double measurement
The 250MHz clock count was subtracted for the plot;
‘Double measurement’ is used to calibrate the LSB, here: LSB=18.2ps;
IODELAY calib: LSB= 18.1ps;
Routing inside the FPGA

## Slide 8: A typical channel in-FPGA calibration
(In 78ps steps)
(TDC Count)
Another channel with five calibration events;
Routing specific linearity (FPGA Clock routing, Carry chain routing)
Event by event repeat,  TDC measurement reolution
* event#1
+ event#2
◊ event#3
○ event#4
X event#5

## Slide 9: A lego plot of in-FPGA calibration
(In 78ps steps)
(TDC Count)
Another channel 1000 calibration events.
The resolution (on the majority of the calibration points) is less than one LSB (<18ps)
(Entries)

## Slide 10: vfTDC status and to do lists
To do lists:
FPGA clock distribution specific calibration, FPGA LOCed routing;
Channel offset calibration (external calibration);
Overall TDC resolution (PCB noise, etc);
Real situation performance and feedback
Status:
192 Channel TDC compiled ;
VME data readout;
Minimum pulse width 3ns, separation 3ns: guaranteed by design;
LSB: 18ps;
In-FPGA calibration resolution: <18ps
Extension:
The Trigger Supervisor board can be used as a 64 channel TDC. The IOs are fully differential on the PCB, and a higher overall resolution can be achieved

## Slide 11: External inputs for overall noise/resolution
8 channels of test signal, synced with ClkVme, 20ns high, 20 ns low.
Various inputs are tried.
Example plots with randomly choosen channels.
8-Ch
ECL
Ouptuts
CAEN
A395A
Ch97-128
Ch65-96
Ch33-64
Ch1-32
Ch129-192
The signal looks like a clock
20ns
20ns
The skew between the eight channels are less than 0.4ns (peak to peak)

## Slide 12: Channel to Channel measurement
8 channels of test signal, synced with ClkVme, 20ns high, 20 ns low.
Various inputs are tried.
Example plots with randomly choosen channels.
The signal looks like a clock. With ~1us DAQ window, there are ~50 edges per event
20ns
20ns
The skew between the eight channels are less than 0.4ns (peak to peak)
There are 35K events for the plots
Timing measurement difference between Ch#n and the first channel (CH#0 and Ch#80)
*RED: Edge=1, rising edge
BLUE: edge = 0, falling edge

## Slide 13: Channel to Channel measurement
Channel#4 as example, further analyze it
For Edge=1, and Fine<35, the DT = 42.4+-1.5
 TDC measurement RMS: 20 ps

## Slide 14: Within the same Channel measurement
Pulse width measurement
Channel#5 as example, plot all the ~50 edges in the right plot
Pulse high: ~20.8ns, pulse low: ~ 19ns
Average: 1277, (1280=20ns);
Fine<50 region: 1239+-1.7 (edge=1)
 TDC measurement RMS: ~25 ps

## Slide 15: Within the same Channel measurement
Some pulses are measured in two consecutive clock cycles, when the carry chain range is longer than the clock period (feature of the carry chain).
Channel#5 as example,
The 2n range corresponds to 128-DeltaT
Fine>128 has different range than that for Fine<128. This means that the Clk250 is not perfectly 50% duty cycle.
Duty cycle difference: 100ps (/4ns)
LSB: 2ns/105=19ps