Front Panel Signal Distribution Module for the FADC250

SD-FP

Ed Jastrzembski

Data Acquisition Group

9/5/08

Summary

The signal distribution module (SD-FP) supports a set of up to seven FADC250 modules through their front panel interfaces.  Common CLOCK, TRIGGER, and SYNC_RESET signals are distributed with low skew to each of the FADC boards.  BUSY signals from the FADC modules are collected and ORed together.  MOD_TRIGGER signals from the FADC modules are collected and driven out for use by external logic.

The SD-FP can be used in stand-alone mode, or together with the Trigger Interface module (TI).  In stand-alone mode, the SD-FP can serve as the source of the 250 MHz CLOCK, or an external CLOCK signal can be accepted.  External TRIGGER and SYNC_RESET signals are then synchronized to the chosen clock (as required by the FADC module).  When in TI mode, the SD-FP simply fans out the CLOCK, TRIGGER, and SYNC_RESET signals provided by the TI.

The SD-FP module can generate TRIGGER and SYNC_RESET pulses by writing to a module register.  These signals are always synchronized to the selected clock (internal or external).  The programmer has access to the current BUSY status of each connected FADC module, and can set a mask identifying the FADC modules that can contribute to the ORed BUSY output.

The SD-FP module also interfaces CLOCK, TRIGGER, SYNC_RESET, MOD_TRIGGER, and BUSY signals onto user pins of the VME P2 connector.  This allows expansion of the system to more than seven FADC modules if a P2 distribution scheme is introduced.

Input/Output Signals

Figure 1 identifies the SD-FP front panel connectors.  Tables 1 - 4 identify the signals carried and their pin assignments.

Table 1.  Connector 1-7 signal definition (FADC  SD-FP).  (Pin 1 is lower-left.)

Signal name 		Direction	Level		Pin # (Q, /Q)

-------- 			--------		--------    	1, 7

BUSY 			input		LVDS		2, 8

MOD TRIGGER	input		LVDS		3, 9

TRIGGER 		output		LVPECL	4, 10

SYNC_RESET 		output		LVPECL	5, 11

CLOCK		output		LVPECL	6, 12

Table 2.  Connector A signal definition (TI  SD-FP).  (Pin 1 is lower-right.)

Signal name 		Direction	Level		Pin # (Q, /Q)

OR BUSY		output		NECL		1, 2

EXT TRIGGER	input		NECL		3, 4

EXT SYNC_RESET 	input		NECL		5, 6

(Ground) 		------		(0 V)		7, 8

EXT CLOCK		input		NECL		9, 10

Table 3.  Connector B-E signal definition.  (B-E: top-to-bottom Lemo connectors.)

Signal name 		Direction	Level		Connector

EXT CLOCK		input		NIM		B

EXT SYNC_RESET 	input		NIM		C

EXT TRIGGER	input		NIM		D

OR BUSY		output		NIM		E

Table 4.  Connector F signal definition (SD-FP => EXT_LOGIC). (Pin 1 is lower-right.)

Signal name 		Direction	Level		Pin # (Q, /Q)

MOD 1 TRIGGER	output		NECL		1, 2

MOD 2 TRIGGER	output		NECL		3, 4

MOD 3 TRIGGER	output		NECL		5, 6

MOD 4 TRIGGER	output		NECL		7, 8

MOD 5 TRIGGER	output		NECL		9, 10

MOD 6 TRIGGER	output		NECL		11, 12

MOD 7 TRIGGER	output		NECL		13, 14

MOD P2 TRIGGER	output		NECL		15, 16

Notes:	(1) MOD TRIGGER (connector ‘n’) => MOD ’n’ TRIGGER (connector F); (n = 1…7).

- Inputs with the same name on connectors A & B-E are ORed together.

Figure 1.  SD-FP front panel

SD-FP Registers

The SD-FP is programmed by the user through VMEbus protocols (ANSI/IEEE STD1014-1987).  The device meets all VMEbus standards.  The VME Trigger Interface is categorized as an A16 - D16 VMEbus slave.   All storage locations can be accessed as both Short Supervisory and Short Non-privileged data.

The base address (A15 – A6) is selected by DIP switches on the board.  An open switch element represents a ‘1’.  The module occupies 64 bytes of VME address space organized in 32 2-byte registers.  Four registers are currently defined.  The remaining space is reserved for testing and future use.

1.  CONTROL/STATUS REGISTER (CSR)	[addr = 0]

(0) BUSY STATUS (R) – state of the BUSY output of the module (‘1’ = BUSY asserted).  Includes the OR of all enabled FADC busy signals (BUSY ENABLE register) and the module soft busy assert bit (CTRL(15)).

(1) – (4) – not used (read as ‘0’)

(5)  PULSE TRIGGER (W) – writing a ‘1’ to this bit generates a pulse on the TRIGGER outputs if software triggers are enabled (CTRL(5) = 1).

(6)  PULSE SYNC_RESET (W) – writing a ‘1’ to this bit generates a pulse on the SYNC_RESET outputs if software sync_resets are enabled (CTRL(9) = 1).

(7)  INITIALIZE (W) – writing a ‘1’ to this bit resets the board to the power-up state.

(15) – (8) – firmware version.

2.  CONTRL (CTRL)	[addr = 2]

(0)  CLOCK SELECT (R/W) – ‘0’ = internal, ‘1’ = external.

(1) – (3) – not used (read as ‘0’)

(4)  TRIGGER SELECT (R/W) – ‘0’ = synchronize TRIGGER input to selected CLOCK,  ‘1’ = DO NOT synchronize TRIGGER to CLOCK.

(5)  SOFT TRIGGER ENABLE (R/W) – ‘1’ = enable, ‘0’ = disable.

(6) – (7) – unused (read as ‘0’)

(8)  SYNC_RESET SELECT (R/W) – ‘0’ = synchronize SYNC_RESET input to selected CLOCK,  ‘1’ = DO NOT synchronize SYNC_RESET to CLOCK.

(9)  SOFT SYNC_RESET ENABLE (R/W) – ‘1’ = enable, ‘0’ = disable.

(10) – (14) – unused (read as ‘0’)

(15) SOFT BUSY ASSERT (R/W) – ‘1’ = assert

3. BUSY ENABLE	[addr = 4]

(0)  ENABLE FADC 1 BUSY (R/W) – ‘1’ = use in OR busy

(1)  ENABLE FADC 2 BUSY (R/W) – ‘1’ = use in OR busy

(2)  ENABLE FADC 3 BUSY (R/W) – ‘1’ = use in OR busy

(3)  ENABLE FADC 4 BUSY (R/W) – ‘1’ = use in OR busy

(4)  ENABLE FADC 5 BUSY (R/W) – ‘1’ = use in OR busy

(5)  ENABLE FADC 6 BUSY (R/W) – ‘1’ = use in OR busy

(6)  ENABLE FADC 7 BUSY (R/W) – ‘1’ = use in OR busy

(7)  ENABLE P2 BUSY (R/W) – ‘1’ = use in OR busy

(8) - (15) - not used (read as ‘1’)

4.  BUSY STATUS	[addr = 6]

(0)  FADC 1 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 1 busy enabled

(1)  FADC 2 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 2 busy enabled

(2)  FADC 3 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 3 busy enabled

(3)  FADC 4 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 4 busy enabled

(4)  FADC 5 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 5 busy enabled

(5)  FADC 6 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 6 busy enabled

(6)  FADC 7 BUSY STATUS (R) – ‘1’ = busy asserted and FADC 7 busy enabled

(7)  P2 BUSY STATUS (R) – ‘1’ = busy asserted and P2 busy enabled

(8)  FADC 1 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(9)  FADC 2 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(10)  FADC 3 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(11)  FADC 4 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(12)  FADC 5 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(13)  FADC 6 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(14)  FADC 7 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

(15)  P2 BUSY STATUS (unmasked) (R) – ‘1’ = busy asserted

Notes:	(1) Power-up state:  CTRL = 0,  BUSY ENABLE = 0.

(CTRL = 0 => internal CLOCK, synchronize TRIGGER & SYNC_RESET to CLOCK, soft TRIGGER & SYNC_RESET generation disabled.)

(2) Disconnected BUSY inputs (connectors 1-7) default to a high (asserted) state.  Use       the BUSY ENABLE register to enable only those inputs that are connected to functioning FADC modules.